Appendix 2. nPOD Paraffin fixation and processing  times using a Sakura TissueTek VIP300 automatic processor.

| Station # | Solution | Conc. | Time | Set Temp. | Vacuum | Mix |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Formalin | 10% | 16:00 | _________ | off | off |
| 2 | EtOH | 70% | hold | _________ | off | off |
| 3 | EtOH | 70% | 0:35 | _________ | on | on |
| 4 | EtOH | 80% | 0:35 | _________ | on | on |
| 5 | EtOH | 95% | 0:35 | _________ | on | on |
| 6 | EtOH | 95% | 0:35 | _________ | on | on |
| 7 | EtOH | 100% | 0:35 | _________ | on | on |
| 8 | EtOH | 100% | 0:35 | _________ | on | on |
| 9 | Xylene | 100% | 0:45 | _________ | on | on |
| 10 | Xylene | 100% | 0:45 | _________ | on | on |
| 11 | Paraffin | _________ | 0:45 | 60` C | on | on |
| 12 | Paraffin | _________ | 0:45 | 60` C | on | on |
| 13 | Paraffin | _________ | 0:45 | 60` C | on | on |
| 14 | Paraffin | _________ | 0:45 | 60` C | on | on |
|  |  | End time (days/hr:min) | 1 or 3 (weekend)/8:30 AM |  |

Hold time duration in 70% ethanol will be dependent on end time.