An Automated Radiosynthesis of [68Ga]Ga-FAPI-46 for Routine Clinical Use

Laurence Morandeau1, Joseph A. Ioppolo1,2, Eva Alvarez de Eulate1, Shifaza Mohamed1, Danica Cullen1, Ali H. Asad1, Roslyn J. Francis2,3,4 Janette Atkinson1

1Radiopharmaceutical Production and Development Centre (RAPID), Medical Technology and Physics Department, Sir Charles Gairdner Hospital, Hospital Avenue, Nedlands, Western Australia, Australia

2National Imaging Facility, Centre for Microscopy, Characterization and Analysis, University of Western Australia, Crawley, Western Australia, Australia

3Department of Nuclear Medicine, Sir Charles Gairdner Hospital, Hospital Avenue, Nedlands, Western Australia, Australia

4Medical School, University of Western Australia, Crawley, Western Australia, Australia

SUPPLEMENTARY MATERIAL:

| Table S1. Reported automated syntheses and quality control details for [68Ga]Ga-FAPI-46. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Reference publications | Spreckelmeyer et al., 2020 (26) | Spreckelmeyer et al., 2020 (26) | Boonkawin & Chotipanich, 2021 (27) | Nader et al., 2022 (28) | Da Pieve et al., 2022 (29) | Alfteimi et al., 2023 (30) | Alfteimi et al., 2023 (30) | Alfteimi et al., 2023 (30) | Plhak et al., 2023 (31) | Rubira et al., 2024 (32) | Mallapura et al., 2023 (33) | Rosenberg et al., 2023 (34) | *This work |
| 68Ga source (generator type or other) | Eckert & Ziegler Gallia Pharm | Eckert & Ziegler Gallia Pharm | ITG | Eckert & Ziegler Gallia Pharm | IRE GalliAd | Eckert & Ziegler Gallia Pharm | Eckert & Ziegler Gallia Pharm | iThemba | IRE GalliAd | IRE GalliAd | Eckert & Ziegler | Cyclotron liquid target 68Zn(p,n)68Ga | Eckert & Ziegler IGG100 |
| Synthesizer | Eckert & Zielger Modular Lab PharmTracer | Eckert & Zielger Modular Lab eazy | ITG iQS-TS | TRASIS EasyOne | TRASIS AiO | TRASIS EasyOne | Synthra | Scintomics GallElut | Scintomics GRP | Elysia Raytest GAIA V2 | iMiDEV microfluidic | GE FASTlab 2 | iPHASE Multisyn |
| Pre-purification SPE cartridge | SCX | SCX | none | none | SCX, 100 mg | none | none | none | SCX | none | Chromafx PS-H+ | ZR resin, Ag1-X8 resin, TK 200 resin | SCX, 100 mg |
| Precursor amount (μg) | 50 | 50 | 50 | 50 | 13 | 50 | 50 | 50 | 50 | 30 | 20 | 50 | 50 |
| Reaction buffer, additives | Sodium ascorbate, ascorbic acid | Sodium ascorbate, ascorbic acid | Sodium acetate | Ascorbic acid | Sodium acetate, sodium ascorbate | Sodium acetate, ascorbic acid | HEPES, ascorbic acid | HEPES, ascorbic acid | HEPES, DMSO | HEPES | Sodium acetate, sodium ascorbate | Sodium acetate, ascorbic acid | Sodium acetate, ascorbic acid |
| Reaction temperature (°C) | 95 | 98 | 95 | 95 | 95 | 90 | 90 | 90 | 125 | 97 | 91 | 95 | 95 |
| Reaction time (min) | 10 | 10 | 10 | 20 | 10 | 4 | 4 | 4 | 6 | 8 | 10 | 10 | 10 |
| Post-purification SPE cartridge | CM | CM | Sep-Pak C18 | Oasis HLB Plus 225 mg | Oasis HLB 30 mg, Sep-Pak QMA Plus Light 130 mg | Oasis HLB | Sep-Pak C18 | Sep-Pak C18 Plus | Sep-Pak Light C18 | Sep-Pak C18 | HLB | Accell Plus CM short | Strata X 60mg |
| Formulation & Stabilisers | Reaction mixture passed through CM, diluted with 0.9% NaCl | Reaction mixture passed through CM, diluted with 0.9% NaCl | 0.5 mL Ethanol, 9.5 mL 0.9% NaCl | 1 mL Ethanol, 9 mL 0.9% NaCl | 0.5 mL 1:1 ethanol water, 1 mg ascorbic acid | 0.9 mL ethanol, 9.1 mL 0.9% NaCl | 3 mL ethanol, 10 mL 0.9% NaCl | 2 mL 1:1 ethanol: water, 10 mL 0.9% NaCl | 2 mL ethanol/water 1:1, 15 mL PBS | 1.5 mL 60% ethanol, 8.6 mL 0.9% NaCl | 1.2 mL 56% ethanol, 8 mL 0.9% NaCl, 5 mg/mL sodium ascorbate | Reaction mixture passed through CM, diluted with 0.9% NaCl | 0.5 mL ethanol, 10.5 mL 0.9% NaCl, 5 mg sodium ascorbate |
| Total procedure time (min) | Not reported | Not reported | 15 | 35 | 35 | Not reported | Not reported | Not reported | 38 | 24 | Not reported | 45 | 25 |
| Yield (MBq) | 740-1700 | 740-1700 | 850-930 | 360-850 | Not reported | Not reported | Not reported | Not reported | 863 ± 129 (n=7) | 370 ± 20 (n=3) | 381 ± 31 (n=3) | > 1480 | 570-1710 (n=6) |
| Yield (%, dc/ndc if reported) | 95.2 ± 1.4 (n=3,) | 89.7 ± 6.7 (n=3) | >70 | 56 ± 5 (ndc), 81 ± 5 (dc) (n=10) | 66.0 ± 7.6 (dc) (n=10) | 92.45 (dc) | 92.86 (dc) | 93.32 (dc) | 72.6 ± 4.9 (dc) (n=7) | 80.5 ± 2.7 (n=3) | 44 ± 5 (dc) (n=3) | Not reported | 89.8 ± 4.8% decay corrected yield (n=6) |
| Radiochemical purity by HPLC (%) | 99.7 | 99.7 | >95 | >99 | >98 | >99.4 | >99.8 | >99.7 | 97.6 ± 0.3 (n=7) | 97.3 ± 0.5 (n=3) | 97.7 ± 0.5% (n=3) | >95 | >99.2 ± 0.7 (n=6) |
| Radiochemical purity by TLC (%) | 99.9 | 99.9 | >95 | >99 | Not reported | Not reported | Not reported | Not reported | 99.8 ± 0.1 (n=7) | 98.1 ± 0.6 (n=3) | 98.2 ± 0.15% (n=3) | >99% | >98.4 ± 1.3 (n=6) |
| Stability: Radiochemical purity (%) | >95 at 3 h | >95 at 3 h | >95 at 4 h | >98 at 4 h | Not reported | >95 at 3 h | >95 at 3 h | >95 at 3 h | 95.3 ± 0.8 at 3 h | ³91.6 at 4 h | 95.8 ± 0.80 at 3 h | >95 at 4 h | 99.1 ± 0.9 at 3 h (HPLC) and 98.0 ± 1.8 (ITLC) |
| FAPI-46 content | Not reported | Not reported | Not reported | < 8 μg/mL | Not reported | Not reported | Not reported | Not reported | 4.1 ± 0.4 µg/mL | Specific activity 12.3 ± 0.6 MBq/μg (n=3) | Molar activity 30 ± 11 GBq/μmol (n=3) | Not reported | Not reported |
| pH | 4.2 | 4.2 | 4.0-8.0 | 7.0 | Not reported | 6.5 | 7.0 | 7.5 | 7.0 | 6 | 5.83 ± 0.1 (n=3) | 4.5 | 6.5 |
| Appearance | Colorless | Colorless | Clear, colorless, particulate free | Colourless | Not reported | Clear & colorless | Clear & colorless | Clear & colorless | Clear & colorless | Clear, colorless | Clear or slightly yellow, free of particles | Clear colorless, no particulates | Clear, colorless, particulate free |
| Endotoxins (EU/mL) | < 5.0 | < 5.0 | <14.58 | < 5.0 | Not reported | < 5.0 | < 5.0 | < 5.0 | 0.05 | Not reported | < 5.0 | < 1.00 | <10.0 |
| Sterility | Sterile | Sterile | Sterile | Sterile | Not reported | Pass | Pass | Pass | Sterile | Not reported | Not reported | No growth | No growth |
| Breakthrough / RadNuc impurity | 0.0002% | 0.0002% | <0.005% | 0.0002% | Not reported | 0.0004% | <0.0001% | <0.0001% | Not reported | Not reported | Not reported | <0.159% 66/67Ga | <0.001% |
| Residual solvents | Not reported | Not reported | <10% ethanol | Not reported | Not reported | Not reported | Not reported | Not reported | 6% ethanol | Not reported | 67,000 ± 2,000 ppm | Not reported | 37,388 ± 4,113 ppm |

Section 1 - Preliminary experiments: Optimization of the reaction time and selection of post-purification SPE cartridge for the automated production of [68Ga]Ga-FAPI-46 on the MS radiosynthesizer.

NOTE: The preliminary experiments were conducted following Steps 1 to 5 described in the protocol, omitting the installation of the post-purification SPE cartridge and the ethanol, water for injection, and saline for injection vials.

1.1.1. Set-up the synthesizer and start synthesis as described in Sections 1-6 of the protocol with the following modifications:

1.1.1.1. Remove and set aside the postpurification SPE cartridge from the kit in Step 3.2.4.6 (Figure 1 A-C, Manifold 2 - valve 4).

1.1.1.2. Pre-condition the four postpurification SPE cartridges selected for this study according to the manufacturer’s instructions (see SPE details in Materials file).

1.1.1.3. Omit step 4.1

1.1.1.4. Omit step 4.2

1.1.1.5. Omit step 4.9

1.1.1.6. Omit step 4.10 and press PAUSE instead of NEXT on user interface.

1.1.1.7. Enter step number corresponding to Ready to elute 68Ga generator on user interface.

1.1.1.8. Press PAUSE to resume on user interface.

1.1.1.9. Press NEXT to start elution.

1.1.1.10. Press PAUSE when the step message Radiolabeling reaction starts and start timer.

1.1.2. To determine the optimal reaction time, draw samples from the reaction mixture at time points 5 min, 15 min and 20 min, following the steps below:

1.1.2.1 Disconnect the side reactor port tubing from the reactor.

1.1.2.2 Disconnect the center port tubing from the reactor.

1.1.2.3. Insert a 1 mL syringe (without needle) onto the center port of the reactor and draw out 50 µL of sample.

1.1.2.4. Transfer the sample to a safe lock 1.5 mL tube and place in a lead pot for analysis.

1.1.2.5. Reconnect the center and side port tubings.

1.1.2.6. Determine the % radiochemical conversion of each time point using analytical radio-TLC analysis by dividing AUC corresponding to [68Ga]Ga-FAPI-46 peak (solvent front) by the sum of the AUC of all radioactive signals (origin + solvent front) in the radio-detector chromatogram, times 100. (Table 2 summarizes the results obtained)

NOTE: During the sampling, extra care was taken to avoid spills of radioactive mixture due to build-up of pressure in the reactor under high temperature.

1.1.3. At the end of the 20 min reaction press PAUSE on user interface to resume synthesis. The synthesizer will automatically switch to the next step message Cooling down reactor.

1.1.4. Allow cooling of the reactor for 270 seconds by monitoring the “elapsed” box.

1.1.5. Press ABORT on user interface to abort the automated production.

NOTE: it is important to either PAUSE or ABORT the synthesis before the end of the 300 seconds cooling step to avoid the synthesizer switching to the next step “trapping production onto SPE cartridge” which would send the reaction mixture to the waste in the absence of a SPE purification cartridge.

1.1.6. For the selection of the optimal SPE purification cartridge, perform manual SPE purifications following the steps below:

1.1.6.1.	 Disconnect the tubing from reactor and close off both ports.

1.1.6.2.	 Transfer reactor in lead pot and measure activity in a dose calibrator.

1.1.6.3.	Manually transfer 6 mL of water for injection to the reactor using a syringe connected to the center port of the reactor.

1.1.6.4.	Aliquot the reaction mixture into four 5 mL syringes (2 mL fraction in each syringe).

1.1.6.5.	Dilute each fraction with 1.3 mL of water for injection.

1.1.6.6.	Measure activity of each fraction decay corrected to time of measurement in step 1.1.6.2.

1.1.6.7.	Transfer each diluted reaction mixture fraction through the designated SPE cartridge.

1.1.6.8.	Rinse each post-purification SPE cartridge with 5 mL of water for injections, flush with 5 mL of air and elute with 1 mL of ethanol into a 1.5 mL tube.

1.1.6.9.	Measure the radioactivity of the recovered fractions (ethanol fraction, waste and SPE) in a dose calibrator for each SPE cartridge and decay correct to time of measurement in step 1.1.6.2.

1.1.6.10. 	Calculate the % recovery by dividing the activity of the ethanol fraction by the sum of all recovered fractions (ethanol fraction, waste and SPE), times 100.

1.1.6.11.	Assess radiochemical identity with analytical radio-HPLC, by verifying the retention times of the [68Ga]Ga-FAPI-46 sample and a non-radioactive [natGa]Ga-FAPI-46 reference standard are similar (allowing for a small delay volume between the radioactivity and UV detectors).

1.1.6.12.	Quantify radiochemical % purity with analytical radio-HPLC using equation (1). Table 3 summarizes the results obtained.

Radiochemical % purity = Area under the curve (AUC) of the [68Ga]Ga-FAPI-46 peak Sum of the AUCs of all radioactive peaks in the radiodetector chromatogram  × 100		(1)

1.1.6.13.	Quantify radiochemical % purity with analytical radio-TLC using equation (2). Table 3 summarizes the results obtained.

Radiochemical % purity = AUC of the [68Ga]Ga-FAPI-46 peak (solvent front)Sum of the AUCs of all radioactive peaks origin+solvent front in the radiodetector chromatogram  × 100  (2)

Figure S1. Image of A) reagents for automated synthesis of [68Ga]Ga-FAPI-46. B) Ancillaries set. C) Photo of final set-up for automated synthesis of [68Ga]Ga-FAPI-46 on MS synthesizer.

Figure S2. UV-chromatogram of 10 mg/L [natGa]Ga-FAPI-46 standard (red solid line – Rt = 4.5 min) and 10 mg/L FAPI-46 precursor (blue solid line – Rt = 4.8 min).

Figure S3. Radio-chromatogram of [68Ga]Ga-FAPI-46 for a high activity (2.8 GBq) radiolabeling reaction without ascorbic acid in the reactor.