## Sheet1
Problem | Causes | Solutions
Poor Crystallization (MALDI-ToF MS target plate preparation) | The cation:analyte:matrix mixture is not homogeneous due to poor solubility. | Test the solubility of the polymer analyte and choose a matrix with similar polarity to the analyte.
| The cation:analyte:matrix mixture contains a much lower concentration of analyte than needed. | Make an analyte sample solution within the range of 1 mg/mL to 10 mg/mL.
The cation:analyte:matrix mixture contains a much higher concentration of analyte than optimal.
| Sample preparation is not optimized | Optimize sample preparation using a 3×3 ratio grid table.
Low Signal-to-Noise (S/N) | The sample preparation mixture is not homogeneous due to poor solubility | Select the appropriate matrix and cation based on polarity or consult literature for recipes used with similar polymers.
| Sample preparation is not optimized. | Optimize sample preparation using a 3×3 ratio grid table.
| Matrix stock solution shelf-life has expired or deactivated due to storage conditions. | Make a new matrix stock solution.
| The laser power is too low to effectively ionize the sample. | Increase the laser power.
| Default MALDI-ToF MS method was not optimized (e.g. for biological samples). | Test alternative acquisition methods or optimize the instrument’s method parameters.
Fragmentation | The laser power is too high (50% or greater), causing the polymer to fragment. | Reduce the laser power to minimize fragmentation.
Matrix Interference | Matrix cluster peaks may interfere with analysis, especially when analyzing low-mass polymers (below 2000 Da). | Perform a blank acquisition using a cation:matrix mixture to identify matrix-related peaks that may also appear in the cation:analyte:matrix sample acquisition.
Lower the intensity of matrix peaks by reducing the matrix concentration in the stock solution (e.g., 10 mg/mL or 5 mg/mL).
Unanticipated Peaks | Metastable peaks from in-source ion decomposition or delayed fragmentation during ion transit through the flight tube. | Perform data acquisition in linear mode; the absence of the suspected peak confirms it was a metastable peak. Metastable peaks typically have poor resolution.
| Fragmentation can occur with a high laser power. | Lower the laser power and expect the intensity of fragmentation to decrease.
| Matrix interference may introduce unexpected signals. | See the solutions listed above under Matrix Interference.
| Unidentifiable peaks may arise from side reactions, contamination, or alternative cationization pathways. | Analyze the synthetic reaction scheme for possible side reactions that produce unexpected end groups.
Investigate potential alternative cationization sources by acquiring spectra using different cationizing agents such as Li⁺, K⁺, and Ag⁺. If the observed m/z values remain unchanged despite the change in cation, the analyte is likely undergoing self-ionization.