Supplemental File 2

- 3D DATA ACQUISITION METHOD FOR TEMPERATURE TRACKING

  - Create a 2-axis temperature data acquisition system using two servos, 3D printed parts, and an electronic development board.

  - Install the system to measure spatial temperature variations inside the chamber with the refrigerator door closed.

Supplemental Figure S1: Design and implementation of the 2-axis temperature data acquisition system. (A) CAD model showing all 3D-printed components and the two servos used in the 2-axis temperature data acquisition system. (B) Assembled mechanical sampling device, with 3D-printed parts (red and grey) mounted on a mock chamber (blue). The device employs two servos to control the movement of a thermocouple within the chamber.

  - Measure and mark 0.57 cm increments on the base of the chamber mount (bottom red piece shown in Supplemental Figure S1B).

  - Write and run an automated script to measure temperature at each point in a 6 × 6 grid with 0.28 cm increments in the XZ plane.

  - After recording temperatures for one 6 × 6 grid, physically move the data logger along the Y-axis.

Supplemental Figure S2: Spatial temperature distribution inside the chamber at 0 °C setpoint. With the temperature chamber set to 0 °C, the internal temperature was sampled five times across 3D space. (A) The distribution of temperatures indicates an average chamber temperature of -1.42 °C. (B-D) The average temperatures along the X, Y, and Z directions, respectively, show no directional trends, indicating spatial temperature uniformity.

With the temperature chamber set to a temperature of 0°C, the temperature in 3D space inside the chamber was sampled five times. (A) The distribution of temperatures in the chamber indicates an average temperature of -1.42°C. (B), (C), (D) The average temperature along the X, Y, and Z directions, respectively, show no trends in temperature with direction.

Repeat data acquisition for each new Y position.

1.7. Perform five trials, varying the order of Y-axis measurements in each trial.

Supplemental Figure S2 shows the aggregate results of sampling the interior chamber space in the X, Y, and Z directions five times. Supplemental Figure S2A indicates that the average temperature is -1.42 °C ± 0.38 °C when the chamber is set to 0 °C. X (Supplemental Figure S2B) was measured from the back to the front of the chamber in 0.28 cm steps; variation was on the order of 10⁻² °C. Y (Supplemental Figure S2C) was measured from left to right in 0.57 cm steps; variation was approximately 0.7 °C. Z (Supplemental Figure S2D) was measured from the center upward in 0.28 cm steps; variation was approximately 0.6 °C. The sampling range was limited to a 1.4 cm × 1.4 cm plane due to mechanical constraints, while the batteries occupied roughly a 2.1 cm × 2.1 cm space. The temperature inside the chamber was generally homogeneous, with hot or cold spots (temperature outliers) detected.

NOTE: The physical constraints of the design of the mechanical sampling device limited the spatial range to a 1.4 cm × 1.4 cm space in the plane, whereas the batteries rest in roughly a 2.1 cm × 2.1 cm space in the center of the chamber. Nonetheless, this test demonstrates that the temperature inside the chamber is generally homogeneous, and hot or cold spots (outliers in temperature) were detected.

- TEMPERATURE vs TIME DATA ACQUISITION METHOD

  - Use a temperature logger and its built-in software to log temperature data from the chamber over 24 hours (Figure 5A).

  - Recognize that this logger is limited in its ability to record over extended durations.

  - Implement a new electronic temperature logging system to measure temperature over 1 week (Figure 5B).

  - Apply Gaussian smoothing to reduce noise in the large dataset.

  - Calculate standard deviations using the raw (unsmoothed) temperature data.

- 3D PRINTING INSTRUCTIONS

  - Print all mounts and the heat sink clamp using 5 vertical shell perimeters and Grid infill at 50%.

  - Print the battery wiring slider using Grid infill at 100%.

  - Print all other STL files using 5 vertical shell perimeters and Grid infill at 15%.