Supplementary protocol: Signals calibration

- PA pressure:

  - Start BDAS

  - In the main menu go to calibration, select the PAP channel.

  - Switch the TAM-A (Transducer/amplifier module for PA pressure) to measure.

  - Open the 3-way stopcock on the connected P75 to air (PA pressure sensor).

  - Press and hold the auto zero button on the TAM-A (PAP) for 3 seconds.

NOTE: The red led lights up and goes off after 3 seconds. On the Bargraph on the TAM-A (PAP), one bar at zero is seen and the digital display on the servo controller for perfusion (SCP) module shows 0. (Display switch on X position, X shows the pressure, Y shows the flow).

  - In BDAS, click on “Start” at lower calibration point, wait 10 seconds and click “Stop” (0 V = 0 cm H2O).

  - Connect the manometer to the open port on the 3-way stopcock. Set it to 10 cm H2O.

NOTE: The bargraph on the TAM-A (PAP) goes up. The display on the SCP shows 10 cm H2O. If not, adjust with the trimmer “CAL DISP X” so that it shows 10 cm H2O. Hold the pressure on the manometer on 10 cm H2O while adjusting with the trimmer.

  - In BDAS click on “Start” at upper calibration point, wait 10 seconds and click “Stop”. A voltage which is around 2V is seen and is equal to 10 cm H2O.

  - Click OK.

  - Switch the 3-way stopcock on the P75 back to the measure position.

  - Proceed the same for LAP calibration.

- Weight calibration:

  - Attach the needle to the weight cell.

  - Set the Edema balance module (EBM) to measure.

  - In the main menu go to calibration, select the weight lung channel.

  - Press the AUTO TARE button on the EBM. It gives the zero weight value to the BDAS software (lower calibration point).

  - Hang the weight on the needle.

  - Do the upper calibration in BDAS at 1000 mg.

  - Click OK.

- Flow calibration:

  - Set the flow at 10 ml/ min.

  - Zero a scale with an empty and clean dish.

  - Collect perfusate at the PA connection level for 1 minute precisely.

  - Weight the dish with the collected perfusate, it should indicate 10 g.

NOTE: If not, adjust with the trimmer “CAL DISP Y” so that it shows the weight measured value (Display switch on Y position, Y shows the flow, X shows the pressure).

  - On BDAS, in the main menu go to calibration, select the flow channel.

  - Click on “Start”. Indicate this weight measured value at upper calibration point, wait 10 seconds and click “Stop.

  - Click OK.

  - Start the DATA acquisition.